## Sheet
7 | SHREE PACETRONIX LTD.
ADDRESS: PLOT NO. 15, SECTOR II, PITHAMPUR, INDORE –454775 (M.P.) INDIA
TITLE: AUTO TESTER ELECTRICAL TESTING FORM FOR DUAL CHAMBER PROMRI CIRCUIT
FORMAT NUMBER – ABCDEF ¦ DEPARTMENT – DEPT. ¦ EFFECTIVE DATE – START DATE ¦ REV – XX
Batch no. | BATCH NO.
Pacemaker Model no. | MODEL NO.
|  |  | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No.
|  | Set Value: | Tolerance: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs:
TEST 1: | AMPLITUDE AND LEAKAGE TEST
| BIPOLAR AMPLITUDE IN UNIPOLAR OUTPUT – A CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| BIPOLAR AMPLITUDE - A CHANNEL | 3.8 |  | 3.72 | 3.56 | 3.56 | 3.64 | 3.56 | 3.48 | 3.64
| BIPOLAR AMPLITUDE - A CHANNEL | 2.5 |  | 2.44 | 2.28 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36
| BIPOLAR AMPLITUDE - A CHANNEL | 1.7 |  | 1.72 | 1.56 | 1.56 | 1.64 | 1.72 | 1.64 | 1.72
| UNIPOLAR AMPLITUDE IN BIPOLAR OUTPUT – A CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| UNIPOLAR AMPLITUDE - A CHANNEL | 3.8 |  | 3.8 | 3.48 | 3.56 | 3.56 | 3.56 | 3.56 | 3.64
| UNIPOLAR AMPLITUDE - A CHANNEL | 2.5 |  | 2.44 | 2.28 | 2.28 | 2.36 | 2.36 | 2.28 | 2.36
| UNIPOLAR AMPLITUDE - A CHANNEL | 1.7 |  | 1.72 | 1.56 | 1.56 | 1.64 | 1.64 | 1.56 | 1.72
| BIPOLAR AMPLITUDE IN UNIPOLAR OUTPUT – V CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| BIPOLAR AMPLITUDE - V CHANNEL | 3.8 |  | 3.76 | 3.44 | 3.44 | 3.44 | 3.44 | 3.44 | 3.6
| BIPOLAR AMPLITUDE - V CHANNEL | 2.5 |  | 2.56 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32
| BIPOLAR AMPLITUDE - V CHANNEL | 1.7 |  | 1.68 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6
| UNIPOLAR AMPLITUDE IN BIPOLAR OUTPUT – V CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| UNIPOLAR AMPLITUDE - V CHANNEL | 3.8 |  | 3.76 | 3.44 | 3.44 | 3.44 | 3.52 | 3.44 | 3.52
| UNIPOLAR AMPLITUDE - V CHANNEL | 2.5 |  | 2.56 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32
| UNIPOLAR AMPLITUDE - V CHANNEL | 1.7 |  | 1.76 | 1.6 | 1.6 | 1.6 | 1.68 | 1.6 | 1.6
| CURRENT | 27 |  | 28.04 | 27.02 | 27.54 | 28.11 | 26.14 | 27.61 | 27.91
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 2: | SENSITIVITY TEST
| BIPOLAR – A CHANNEL | 0.6 |  | 0.64 | 0.61 | 0.7 | 0.61 | 0.66 | 0.6 | 0.68
| BIPOLAR – A CHANNEL | 1 |  | 1.15 | 1.1 | 1.15 | 1.15 | 1.2 | 1.1 | 1.15
| BIPOLAR – A CHANNEL | 1.4 |  | 1.55 | 1.5 | 1.55 | 1.65 | 1.5 | 1.35 | 1.6
| UNIPOLAR – A CHANNEL | 0.6 |  | 0.62 | 0.62 | 0.71 | 0.69 | 0.61 | 0.59 | 0.74
| UNIPOLAR – A CHANNEL | 1 |  | 1.1 | 1.15 | 1.15 | 1.15 | 1.15 | 1.05 | 1.15
| UNIPOLAR – A CHANNEL | 1.4 |  | 1.5 | 1.55 | 1.55 | 1.65 | 1.6 | 1.5 | 1.6
| BIPOLAR – V CHANNEL | 1.5 |  | 1.4 | 1.3 | 1.3 | 1.4 | 1.3 | 1.4 | 1.3
| BIPOLAR – V CHANNEL | 2.5 |  | 2.4 | 2.2 | 2.2 | 2.3 | 2.4 | 2.3 | 2.2
| BIPOLAR – V CHANNEL | 3.5 |  | 3.2 | 3.1 | 3.1 | 3.4 | 3.3 | 3.2 | 3.1
| UNIPOLAR – V CHANNEL | 1.5 |  | 1.5 | 1.3 | 1.3 | 1.5 | 1.5 | 1.4 | 1.3
| UNIPOLAR – V CHANNEL | 2.5 |  | 2.4 | 2.3 | 2.2 | 2.5 | 2.4 | 2.3 | 2.2
| UNIPOLAR – V CHANNEL | 3.5 |  | 3.3 | 3.1 | 3.1 | 3.4 | 3.3 | 3.1 | 3.1
| CURRENT | 27 |  | 28.65 | 27.82 | 28.6 | 28.68 | 27.77 | 28.37 | 28.46
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 3: | URGENT TEST - UNIPOLAR
| URGENT AMPLITUDE - UNIPOLAR - A CHANNEL | 5 |  | 4.92 | 4.6 | 4.6 | 4.76 | 4.68 | 4.68 | 4.84
| URGENT AMPLITUDE - UNIPOLAR - V CHANNEL | 5 |  | 4.88 | 4.56 | 4.72 | 4.8 | 4.64 | 4.72 | 4.72
| URGENT RATE | 60 |  | 60 | 60 | 60 | 60 | 60 | 60 | 60
| URGENT PULSE WIDTH - A CHANNEL | 0.5 |  | 0.41 | 0.41 | 0.41 | 0.4 | 0.4 | 0.41 | 0.41
| URGENT PULSE WIDTH - V CHANNEL | 0.5 |  | 0.41 | 0.41 | 0.41 | 0.41 | 0.41 | 0.41 | 0.41
| URGENT CURRENT | 27 |  | 27.94 | 27.19 | 27.67 | 28.12 | 27.32 | 27.9 | 27.94
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 4: | RATE TEST
| BIPOLAR RATE | 50 |  | 49.59 | 49.59 | 49.59 | 49.59 | 49.59 | 49.59 | 49.59
| LOW ENERGY IMPEDANCE - BIPOLAR - A CHANNEL | 500 |  | 487 | 459 | 517 | 533 | 459 | 447 | 502
| LOW ENERGY IMPEDANCE - BIPOLAR - V CHANNEL | 500 |  | 547 | 486 | 531 | 564 | 500 | 486 | 500
| UNIPOLAR RATE | 50 |  | 49.59 | 49.59 | 49.59 | 49.59 | 49.59 | 49.59 | 49.59
| LOW ENERGY IMPEDANCE - UNIPOLAR - A CHANNEL | 500 |  | 487 | 473 | 487 | 517 | 459 | 459 | 502
| LOW ENERGY IMPEDANCE - UNIPOLAR - V CHANNEL | 500 |  | 564 | 486 | 531 | 564 | 486 | 472 | 500
| CURRENT(LOW ENERGY) | 13 |  | 11.12 | 11.18 | 11.41 | 11.33 | 10.96 | 11.59 | 11.25
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 5: | PULSE WIDTH TEST
| BIPOLAR PULSE WIDTH - A CHANNEL | 0.3 |  | 0.28 | 0.28 | 0.29 | 0.28 | 0.28 | 0.29 | 0.29
| BIPOLAR PULSE WIDTH - A CHANNEL | 1.2 |  | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
| BIPOLAR PULSE WIDTH - A CHANNEL | 1.5 |  | 1.51 | 1.51 | 1.51 | 1.5 | 1.5 | 1.51 | 1.51
| UNIPOLAR PULSE WIDTH - A CHANNEL | 0.3 |  | 0.29 | 0.28 | 0.29 | 0.28 | 0.28 | 0.29 | 0.29
| UNIPOLAR PULSE WIDTH - A CHANNEL | 1.2 |  | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
| UNIPOLAR PULSE WIDTH - A CHANNEL | 1.5 |  | 1.51 | 1.5 | 1.5 | 1.51 | 1.5 | 1.51 | 1.51
| BIPOLAR PULSE WIDTH - V CHANNEL | 0.3 |  | 0.29 | 0.29 | 0.29 | 0.29 | 0.28 | 0.29 | 0.29
| BIPOLAR PULSE WIDTH - V CHANNEL | 1.2 |  | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.21
| BIPOLAR PULSE WIDTH - V CHANNEL | 1.5 |  | 1.51 | 1.51 | 1.51 | 1.51 | 1.51 | 1.51 | 1.51
| UNIPOLAR PULSE WIDTH - V CHANNEL | 0.3 |  | 0.29 | 0.29 | 0.29 | 0.29 | 0.28 | 0.29 | 0.29
| UNIPOLAR PULSE WIDTH - V CHANNEL | 1.2 |  | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2
| UNIPOLAR PULSE WIDTH - V CHANNEL | 1.5 |  | 1.51 | 1.51 | 1.51 | 1.51 | 1.51 | 1.51 | 1.51
| CURRENT | 27 |  | 27.58 | 27.4 | 27.8 | 27.64 | 27.08 | 27.5 | 27.5
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 6: | IMPEDANCE TEST
| BIPOLAR IMPEDANCE - A CHANNEL | 350 |  | 344 | 336 | 344 | 361 | 327 | 327 | 344
| BIPOLAR IMPEDANCE - A CHANNEL | 1000 |  | 988 | 898 | 1038 | 1094 | 898 | 859 | 1038
| BIPOLAR IMPEDANCE - A CHANNEL | 500 |  | 502 | 487 | 502 | 517 | 459 | 447 | 502
| BIPOLAR IMPEDANCE - A CHANNEL | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| BIPOLAR IMPEDANCE - A CHANNEL | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| UNIPOLAR IMPEDANCE - A CHANNEL | 350 |  | 344 | 327 | 352 | 361 | 327 | 320 | 344
| UNIPOLAR IMPEDANCE - A CHANNEL | 1000 |  | 988 | 941 | 1038 | 1094 | 859 | 859 | 988
| UNIPOLAR IMPEDANCE - A CHANNEL | 500 |  | 502 | 473 | 487 | 517 | 459 | 447 | 502
| UNIPOLAR IMPEDANCE - A CHANNEL | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| UNIPOLAR IMPEDANCE - A CHANNEL | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| BIPOLAR IMPEDANCE - V CHANNEL | 350 |  | 382 | 355 | 392 | 392 | 346 | 346 | 355
| BIPOLAR IMPEDANCE - V CHANNEL | 1000 |  | 1157 | 950 | 1044 | 1157 | 950 | 950 | 995
| BIPOLAR IMPEDANCE - V CHANNEL | 500 |  | 547 | 486 | 547 | 564 | 486 | 500 | 500
| BIPOLAR IMPEDANCE - V CHANNEL | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| BIPOLAR IMPEDANCE - V CHANNEL | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| UNIPOLAR IMPEDANCE - V CHANNEL | 350 |  | 392 | 355 | 392 | 392 | 346 | 346 | 346
| UNIPOLAR IMPEDANCE - V CHANNEL | 1000 |  | 1157 | 950 | 1044 | 1157 | 950 | 950 | 995
| UNIPOLAR IMPEDANCE - V CHANNEL | 500 |  | 547 | 486 | 547 | 564 | 472 | 500 | 500
| UNIPOLAR IMPEDANCE - V CHANNEL | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| UNIPOLAR IMPEDANCE - V CHANNEL | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| CURRENT | 27 |  | 27.9 | 27.12 | 28.4 | 28.18 | 28.18 | 27.93 | 28.17
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 7: | BATTERY TEST
| EOL | BATTERY VOLT EOL |  | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL
| ERI STATUS DURING EOL | ERI/EOL/BOL |  | EOL | EOL | EOL | EOL | EOL | EOL | EOL
| ERI | BATTERY VOLT ERI |  | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI
| EOL STATUS DURING ERI | ERI/EOL/BOL |  | ERI | ERI | ERI | ERI | ERI | ERI | ERI
| BATTERY VOLTAGE (BOL) | BATTERY VOLT BOL |  | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL
| ERI/EOL STATUS DURING BOL | ERI/EOL/BOL |  | BOL | BOL | BOL | BOL | BOL | BOL | BOL
| MODE DURING EOL - VVI | VVI |  | VVI | VVI | VVI | VVI | VVI | VVI | VVI
| CURRENT - BOL | 27 |  | 28.74 | 27.4 | 27.85 | 28.74 | 27.5 | 28.17 | 28.64
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 8: | MAGNET TEST
| AMPLITUDE BEFORE MAGNET TEST - A CHANNEL | 3 |  | 2.92 | 2.76 | 2.84 | 2.92 | 2.84 | 2.84 | 2.84
| AMPLITUDE BEFORE MAGNET TEST - V CHANNEL | 3 |  | 2.96 | 2.72 | 2.8 | 2.72 | 2.8 | 2.72 | 2.8
| RATE BEFORE MAGNET TEST | 108 |  | 107.47 | 107.47 | 107.47 | 107.46 | 107.47 | 107.47 | 107.47
| MAGNET MODE - DOO | DOO |  | DOO | DOO | DOO | DOO | DOO | DOO | DOO
| AMPLITUDE DURING MAGNET TEST - A CHANNEL | 4.5 |  | 4.36 | 4.12 | 4.28 | 4.28 | 4.28 | 4.2 | 4.36
| AMPLITUDE DURING MAGNET TEST - V CHANNEL | 4.5 |  | 4.4 | 4.16 | 4.24 | 4.32 | 4.32 | 4.24 | 4.24
| CURRENT DURING MAGNET TEST | 35 |  | 37.49 | 36.51 | 36.73 | 37.4 | 36.74 | 36.99 | 37.25
| RATE DURING MAGNET TEST (BOL) | 96 |  | 95.57 | 95.57 | 95.57 | 95.57 | 95.57 | 95.57 | 95.58
| RATE ERI MAGNET | 83 |  | 83.16 | 83.16 | 83.15 | 83.15 | 83.15 | 83.16 | 83.16
| RATE EOL MAGNET | 64 |  | 63.91 | 63.91 | 63.91 | 63.9 | 63.91 | 63.91 | 63.91
| AMPLITUDE AFTER MAGNET TEST - A CHANNEL | 3 |  | 3 | 2.76 | 2.84 | 2.92 | 2.84 | 2.76 | 2.84
| AMPLITUDE AFTER MAGNET TEST - V CHANNEL | 3 |  | 2.88 | 2.72 | 2.8 | 2.8 | 2.8 | 2.72 | 2.88
| RATE AFTER MAGNET TEST | 108 |  | 107.47 | 107.47 | 107.47 | 107.46 | 107.47 | 107.47 | 107.47
| MAGNET TEST (DOO) - A CHANNEL | PACE COUNT |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| MAGNET TEST (DOO) - V CHANNEL | PACE COUNT |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| CURRENT | 27 |  | 28 | 27.28 | 27.53 | 28.19 | 28.26 | 27.97 | 28.04
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 9: | NOISE TEST
| 11HZ BURST PERIOD TEST - BIPOLAR – A CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - BIPOLAR – A CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11HZ BURST PERIOD TEST - UNIPOLAR – A CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - UNIPOLAR – A CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11HZ BURST PERIOD TEST - BIPOLAR – V CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - BIPOLAR – V CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11HZ BURST PERIOD TEST - UNIPOLAR – V CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - UNIPOLAR – V CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| CURRENT | 27 |  | 27.85 | 28.05 | 27.42 | 27.95 | 28.43 | 28.87 | 27.88
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 10: | HIGH ENERGY TEST
| RATE | 120 |  | 119.1 | 119.1 | 119.09 | 119.08 | 119.08 | 119.08 | 119.1
| AMPLITUDE - A CHANNEL | 7.5 |  | 6.44 | 6.36 | 6.44 | 6.52 | 6.52 | 6.44 | 6.44
| AMPLITUDE - V CHANNEL | 7.5 |  | 6.56 | 6.4 | 6.56 | 6.56 | 6.56 | 6.48 | 6.56
| PULSE WIDTH - A CHANNEL | 1.0 |  | 1.02 | 1.02 | 1.02 | 1.01 | 1.01 | 1.02 | 1.02
| PULSE WIDTH - V CHANNEL | 1.0 |  | 1.02 | 1.02 | 1.02 | 1.02 | 1.02 | 1.02 | 1.02
| HIGH ENERGY IMPEDANCE - A CHANNEL | 500 |  | 517 | 473 | 502 | 517 | 459 | 459 | 502
| HIGH ENERGY IMPEDANCE - V CHANNEL | 500 |  | 564 | 500 | 547 | 582 | 486 | 500 | 515
| CURRENT | 180 |  | 178.76 | 177.27 | 178.61 | 177.68 | 178.67 | 178.46 | 179.2
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 11: | TRAILING EDGE AND SLEW RATE TEST
| LEADING AMPLITUDE - 5V - DDD - BIPOLAR - A CHANNEL | 5.3 |  | 5.16 | 5.08 | 5.08 | 5.08 | 5 | 5 | 5.16
| TRAILING AMPLITUDE - 5V - DDD - BIPOLAR - A CHANNEL | 4.3 |  | 4.44 | 4.28 | 4.44 | 4.44 | 4.36 | 4.36 | 4.52
| AVERAGE AMPLITUDE - 5V - DDD - BIPOLAR - A CHANNEL | 4.8 |  | 4.8 | 4.68 | 4.76 | 4.76 | 4.68 | 4.68 | 4.84
| SLEW RATE - 5V - DDD - BIPOLAR - A CHANNEL | 0.5-3.0 |  | 2.4 | 1.6 | 2 | 1.6 | 2 | 1.2 | 2
| LEADING AMPLITUDE - 5V - DDD - UNIPOLAR - A CHANNEL | 5.3 |  | 5.16 | 5 | 5.16 | 5.16 | 5 | 4.92 | 5.24
| TRAILING AMPLITUDE - 5V - DDD - UNIPOLAR - A CHANNEL | 4.3 |  | 4.52 | 4.36 | 4.36 | 4.44 | 4.36 | 4.36 | 4.52
| AVERAGE AMPLITUDE - 5V - DDD - UNIPOLAR - A CHANNEL | 4.8 |  | 4.84 | 4.68 | 4.76 | 4.8 | 4.68 | 4.64 | 4.88
| SLEW RATE - 5V - DDD - UNIPOLAR - A CHANNEL | 0.5-3.0 |  | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 2 | 2
| LEADING AMPLITUDE - 5V - DDD - BIPOLAR - V CHANNEL | 5.3 |  | 5.2 | 4.96 | 4.96 | 5.12 | 5.04 | 4.96 | 5.04
| TRAILING AMPLITUDE - 5V - DDD - BIPOLAR - V CHANNEL | 4.3 |  | 4.48 | 4.32 | 4.32 | 4.4 | 4.4 | 4.24 | 4.4
| AVERAGE AMPLITUDE - 5V - DDD - BIPOLAR - V CHANNEL | 4.8 |  | 4.84 | 4.64 | 4.64 | 4.76 | 4.72 | 4.6 | 4.72
| SLEW RATE - 5V - DDD - BIPOLAR - V CHANNEL | 0.5-3.0 |  | 1.6 | 1.2 | 0.8 | 2 | 1.2 | 1.2 | 1.6
| LEADING AMPLITUDE - 5V - DDD - UNIPOLAR - V CHANNEL | 5.3 |  | 5.2 | 4.96 | 5.04 | 5.04 | 5.04 | 4.96 | 5.12
| TRAILING AMPLITUDE - 5V - DDD - UNIPOLAR - V CHANNEL | 4.3 |  | 4.48 | 4.32 | 4.4 | 4.48 | 4.24 | 4.4 | 4.4
| AVERAGE AMPLITUDE - 5V - DDD - UNIPOLAR - V CHANNEL | 4.8 |  | 4.84 | 4.64 | 4.72 | 4.76 | 4.64 | 4.68 | 4.76
| SLEW RATE - 5V - DDD - UNIPOLAR - V CHANNEL | 0.5-3.0 |  | 1.6 | 0.8 | 1.6 | 2 | 1.2 | 1.6 | 2
| CURRENT | 27 |  | 27.94 | 27.82 | 27.72 | 28.14 | 27.34 | 28.98 | 28.58
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 12: | RATE RESPONSIVE RATE TEST
| DDDR MODE CURRENT | 33 |  | 33.4 | 31.85 | 33.46 | 32.52 | 32.55 | 32.49 | 32.31
| RATE BEFORE ACTIVITY | 60 |  | 58.82 | 58.82 | 58.82 | 58.82 | 58.82 | 58.82 | 58.82
| RATE CHANGE AFTER ACTIVITY |  |  | 88.71 | 88.45 | 72.02 | 79.46 | 88.44 | 82.45 | 79.48
| R MODE CURRENT AFTER ACTIVITY | 45 |  | 45.01 | 46.13 | 37.64 | 41.76 | 45.73 | 42 | 39.46
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 13: | AV DELAY TEST
| PACE DELAY | 140 |  | 134.62 | 134.61 | 134.61 | 134.68 | 134.65 | 134.65 | 134.62
| PACE DELAY | 245 |  | 243.99 | 243.99 | 243.99 | 244.05 | 244.02 | 244.02 | 243.99
| PACE DELAY | 350 |  | 345.56 | 345.56 | 345.56 | 345.62 | 345.59 | 345.59 | 345.56
| SENSE DELAY | 140 |  | 129.18 | 129.21 | 129.2 | 129.77 | 129.21 | 129.74 | 129.1
| UPPER RATE AV FOLLOW TEST-490MS BURST PERIOD | 150-800 |  | 160.62 | 160.49 | 160.62 | 160.67 | 160.55 | 160.79 | 160.95
| CURRENT | 27 |  | 27.94 | 27.19 | 27.69 | 28.15 | 27.34 | 27.89 | 27.96
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TESTED BY:
NAME:
SIGNATURE: