## Sheet
9
SHREE PACETRONIX LTD.
ADDRESS: PLOT NO. 15, SECTOR II, PITHAMPUR, INDORE –454775 (M.P.) INDIA
TITLE: AUTO TESTER ELECTRICAL TESTING FORM FOR DUAL CHAMBER PROMRI PACEMAKER
FORMAT NUMBER – ABCDEF ¦ DEPARTMENT – DEPT. ¦ EFFECTIVE DATE – Date of start¦ REV – XX
Batch no. | Batch no.
Pacemaker Model no. | Model No. | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage
|  | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No.
| Set Value: | Tolerance: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs:
Found OOO Mode: | OK/FAIL |  | OK | OK | OK | OK | OK | OK | OK | OK | OK
Battery Test | Battery Volt |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
MODE | DDD |  | DDD | DDD | DDD | DDD | DDD | DDD | DDD | DDD | DDD
Bipolar Amplitude in Bipolar O/P-DDD-A CHANNEL | 2.5 |  | 2.28 | 2.52 | 2.36 | 2.28 | 2.44 | 2.36 | 2.44 | 2.52 | 2.44
Bipolar Amplitude in Unipolar O/P-DDD-A CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bipolar Amplitude in Bipolar O/P-DDD-V CHANNEL | 5 |  | 4.56 | 4.56 | 4.56 | 4.4 | 4.48 | 4.48 | 4.48 | 4.48 | 4.56
Bipolar Amplitude in Unipolar O/P-DDD-V CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unipolar Amplitude in Unipolar O/P-DDD-A CHANNEL | 2.5 |  | 2.44 | 2.36 | 2.28 | 2.28 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36
Unipolar Amplitude in Bipolar O/P-DDD-A CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unipolar Amplitude in Unipolar O/P-DDD-V CHANNEL | 5 |  | 4.64 | 4.72 | 4.64 | 4.64 | 4.72 | 4.64 | 4.64 | 4.56 | 4.64
Unipolar Amplitude in Bipolar O/P-DDD-V CHANNEL | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rate Bipolar-DDD | 60 |  | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60
Impedance Bipolar-DDD-A CHANNEL | 500 |  | 550 | 517 | 517 | 459 | 487 | 487 | 517 | 487 | 517
Impedance Bipolar-DDD-V CHANNEL | 500 |  | 582 | 582 | 601 | 500 | 564 | 531 | 547 | 515 | 564
Pulse Width Test-DDD-A CHANNEL | 0.4 |  | 0.41 | 0.41 | 0.4 | 0.41 | 0.41 | 0.41 | 0.41 | 0.41 | 0.4
Pulse Width Test-DDD-V CHANNEL | 0.4 |  | 0.41 | 0.41 | 0.4 | 0.41 | 0.41 | 0.41 | 0.41 | 0.41 | 0.4
Magnet Test Amplitude-A CHANNEL | 4.5 |  | 4.2 | 4.12 | 4.12 | 4.04 | 4.12 | 4.12 | 4.12 | 4.12 | 4.04
Magnet Test Amplitude-V CHANNEL | 4.5 |  | 4.08 | 4.16 | 4.08 | 4 | 4.16 | 4.08 | 4.08 | 4.08 | 4.08
Magnet Test Rate | 96 |  | 95.7 | 95.58 | 95.7 | 95.7 | 95.7 | 95.7 | 95.7 | 95.7 | 95.7
Magnet Test Sensing-DOO | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Amplitude after Magnet-DDD-A CHANNEL | 2.5 |  | 2.36 | 2.76 | 2.52 | 2.68 | 2.76 | 2.6 | 2.68 | 2.6 | 2.6
Amplitude after Magnet-DDD-V CHANNEL | 5 |  | 4.56 | 4.56 | 4.48 | 4.48 | 4.64 | 4.56 | 4.4 | 4.48 | 4.56
Noise Test: Count at 11Hz-DDD-A CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Noise Test: Count at 5Hz-DDD-A CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Noise Test: Count at 11Hz-DDD-V CHANNEL | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Noise Test: Count at 5Hz-DDD-V CHANNEL | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sensitivity Unipolar-DDD-A CHANNEL | 0.6 |  | 0.64 | 0.67 | 0.68 | 0.66 | 0.64 | 0.68 | 0.7 | 0.63 | 0.67
Sensitivity Bipolar-DDD-A CHANNEL | 0.6 |  | 0.68 | 0.69 | 0.75 | 0.7 | 0.66 | 0.73 | 0.7 | 0.64 | 0.69
Sensitivity Bipolar-DDD-A CHANNEL | 1 |  | 1.14 | 1.1 | 1.12 | 1.09 | 1.15 | 1.17 | 1.17 | 1.07 | 1.14
Sensitivity Unipolar-DDD-V CHANNEL | 1.5 |  | 1.6 | 1.5 | 1.5 | 1.6 | 1.5 | 1.4 | 1.4 | 1.5 | 1.5
Sensitivity Bipolar-DDD-V CHANNEL | 1.5 |  | 1.6 | 1.6 | 1.4 | 1.6 | 1.5 | 1.4 | 1.4 | 1.6 | 1.5
Sensitivity Bipolar-DDD-V CHANNEL | 2.5 |  | 2.6 | 2.6 | 2.4 | 2.6 | 2.5 | 2.4 | 2.4 | 2.5 | 2.5
POST AV DELAY - PACE DELAY | 140 |  | 134.46 | 134.62 | 134.68 | 134.8 | 134.65 | 134.61 | 134.68 | 134.61 | 134.68
POST AV DELAY - SENSE DELAY | 125 |  | 129.66 | 129.58 | 129.78 | 129.38 | 129.35 | 129.42 | 129.71 | 129.62 | 129.51
UPPER RATE AV FOLLOW TEST-490ms Burst Period | 150-800 |  | 270.13 | 270.11 | 270.31 | 270.11 | 270 | 270.06 | 270.03 | 270.18 | 270.19
Leading Amplitude 5V -DDD-A CHANNEL | 5.1 |  | 4.92 | 4.84 | 4.92 | 4.84 | 4.92 | 4.92 | 4.84 | 4.84 | 4.92
Trailing Amplitude 5V -DDD-A CHANNEL | 4.4 |  | 4.28 | 4.44 | 4.36 | 4.2 | 4.36 | 4.28 | 4.28 | 4.28 | 4.36
Average Amplitude -DDD-A CHANNEL | 4.8 |  | 4.6 | 4.64 | 4.64 | 4.52 | 4.64 | 4.6 | 4.56 | 4.56 | 4.64
Slew Rate-DDD-A CHANNEL | 0.5-3.0 |  | 1.6 | 2 | 0.8 | 1.6 | 2 | 1.6 | 1.6 | 1.2 | 1.6
Leading Amplitude 5V -DDD-V CHANNEL | 5.1 |  | 4.96 | 4.96 | 4.88 | 4.8 | 4.8 | 4.8 | 4.8 | 4.8 | 4.88
Trailing Amplitude 5V -DDD-V CHANNEL | 4.4 |  | 4.32 | 4.32 | 4.24 | 4.24 | 4.32 | 4.16 | 4.16 | 4.16 | 4.24
Average Amplitude -DDD-V CHANNEL | 4.8 |  | 4.64 | 4.64 | 4.56 | 4.52 | 4.56 | 4.48 | 4.48 | 4.48 | 4.56
Slew Rate-DDD-V CHANNEL | 0.5-3.0 |  | 1.6 | 1.2 | 1.6 | 0.8 | 1.2 | 0.8 | 1.2 | 1.6 | 1.6
Rate Responsive Rate test-DDDR | >60 |  | 87.81 | 85.75 | 83.91 | 85.83 | 89.8 | 104.55 | 101.1 | 104.55 | 85.85
Battery Test | Battery Volt |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
OOO mode check | OK/FAIL |  | OK | OK | OK | OK | OK | OK | OK | OK | OK
| Final Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Tested By
Date
Name
Signature