## Sheet
12
SHREE PACETRONIX LTD.
ADDRESS: PLOT NO. 15, SECTOR II, PITHAMPUR, INDORE –454775 (M.P.) INDIA
TITLE: AUTO TESTER ELECTRICAL TESTING FORM FOR SINGLE CHAMBER WITH R MODE PROMRI PACEMAKER
FORMAT NUMBER – ABCDEF ¦ DEPARTMENT – DEPT. ¦ EFFECTIVE DATE – Date of start¦ REV – XX
Batch no. | Batch no.
Pacemaker Model no. | Model No. | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage | Stage
|  | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No.
| Set Value: | Tolerance: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs:
Found OOO Mode: | OK/FAIL |  | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK
Battery Test | Battery Volt |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
MODE | VVI |  | VVI | VVI | VVI | VVI | VVI | VVI | VVI | VVI | VVI | VVI | VVI | VVI
Bipolar Amplitude in Bipolar mode-VVI | 3.6 |  | 3.44 | 3.28 | 3.36 | 3.36 | 3.44 | 3.28 | 3.36 | 3.36 | 3.6 | 3.6 | 3.52 | 3.6
Bipolar Amplitude in Unipolar mode-VVI | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unipolar Amplitude in Unipolar mode-VVI | 3.6 |  | 3.52 | 3.44 | 3.44 | 3.6 | 3.6 | 3.44 | 3.52 | 3.44 | 3.44 | 3.52 | 3.52 | 3.52
Unipolar Amplitude in Bipolar mode-VVI | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rate Bipolar-VVI | 70 |  | 69.69 | 69.75 | 69.7 | 69.71 | 69.71 | 69.78 | 69.74 | 69.8 | 69.77 | 69.8 | 69.71 | 69.68
Sensitivity Unipolar-VVI | 1.5 |  | 1.8 | 1.9 | 2 | 1.7 | 2 | 1.6 | 1.8 | 1.8 | 1.9 | 2 | 1.8 | 1.7
Sensitivity Bipolar-VVI | 1.5 |  | 1.8 | 1.9 | 1.9 | 1.7 | 2 | 1.6 | 1.8 | 1.8 | 2 | 2 | 1.9 | 1.7
Sensitivity Bipolar-VVI | 2.5 |  | 2.5 | 2.6 | 2.8 | 2.5 | 2.9 | 2.25 | 2.5 | 2.5 | 2.8 | 2.7 | 2.6 | 2.4
Impedance Bipolar-VVI | 500 |  | 549 | 549 | 556 | 549 | 556 | 556 | 527 | 549 | 541 | 534 | 541 | 534
Pulse Width Test-VVI | 0.4 |  | 0.39 | 0.42 | 0.39 | 0.39 | 0.39 | 0.39 | 0.4 | 0.4 | 0.39 | 0.4 | 0.39 | 0.42
Magnet Test Amplitude | 4.5 |  | 4 | 4 | 4.08 | 4 | 4 | 4.08 | 4.08 | 4 | 4.24 | 4.24 | 4.16 | 4.24
Magnet Test Rate | 96 |  | 95.82 | 95.93 | 95.82 | 95.85 | 95.85 | 95.97 | 95.9 | 96.02 | 95.97 | 95.94 | 95.85 | 95.87
Magnet Test Sensing-VOO | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Amplitude after Magnet-VVI | 3.6 |  | 3.36 | 3.36 | 3.36 | 3.28 | 3.36 | 3.28 | 3.28 | 3.28 | 3.44 | 3.52 | 3.44 | 3.52
Noise Test: Count at 5Hz-VVI | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Noise Test: Count at 11Hz-VVI | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Leading Amplitude 5V VVI | 5.1 |  | 5.12 | 5.04 | 5.04 | 5.04 | 5.04 | 4.96 | 5.12 | 5.04 | 5.28 | 5.28 | 5.28 | 5.28
Trailing Amplitude 5V VVI | 4.4 |  | 4.24 | 4.24 | 4.24 | 4.16 | 4.24 | 4.08 | 4.16 | 4.08 | 4.48 | 4.32 | 4.24 | 4.56
Average Amplitude-5-VVI | 4.8 |  | 4.68 | 4.64 | 4.64 | 4.6 | 4.64 | 4.52 | 4.64 | 4.56 | 4.88 | 4.8 | 4.76 | 4.92
Slew Rate-VVI | 0.5-3.0 |  | 2 | 2 | 1.6 | 0.8 | 1.6 | 2 | 1.6 | 2.4 | 1.2 | 2 | 2 | 1.6
Bipolar Amplitude 7.5-VVI | 7.5 |  | 6.88 | 6.8 | 6.88 | 6.96 | 6.88 | 6.96 | 6.96 | 6.96 | 7.2 | 7.28 | 7.2 | 7.36
Rate Responsive Rate test-VVIR | >70 |  | 90.51 | 98.86 | 96.21 | 98.14 | 91.17 | 107.93 | 90.35 | 98.87 | 94.05 | 91.24 | 93.94 | 96.78
Battery Test | Battery Volt |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
OOO mode check | OK/FAIL |  | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK
| Final Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Tested By
Date
Name
Signature