## Sheet
3 | SHREE PACETRONIX LTD.
ADDRESS: PLOT NO. 15, SECTOR II, PITHAMPUR, INDORE –454775 (M.P.) INDIA
TITLE: AUTO TESTER ELECTRICAL TESTING FORM FOR SINGLE CHAMBER NON R MODE PROMRI CIRCUIT
FORMAT NUMBER – ABCDEF ¦ DEPARTMENT – DEPT. ¦ EFFECTIVE DATE – START DATE ¦ REV – XX
Batch no. | BATCH NO.
Pacemaker Model no. | MODEL NO.
|  |  | Serial No. | Serial No. | Serial No. | Serial No.
|  | Set Value: | Tolerance: | Obs: | Obs: | Obs:
TEST 1: | AMPLITUDE AND LEAKAGE TEST
| BIPOLAR AMPLITUDE IN UNIPOLAR OUTPUT | 0 |  | 0 | 0 | 0
| BIPOLAR AMPLITUDE | 3.8 |  | 3.6 | 3.68 | 3.6
| BIPOLAR AMPLITUDE | 2.5 |  | 2.48 | 2.48 | 2.56
| BIPOLAR AMPLITUDE | 1.7 |  | 1.84 | 1.84 | 1.68
| UNIPOLAR AMPLITUDE IN BIPOLAR OUTPUT | 0 |  | 0 | 0 | 0
| UNIPOLAR AMPLITUDE | 3.8 |  | 3.52 | 3.52 | 3.52
| UNIPOLAR AMPLITUDE | 2.5 |  | 2.64 | 2.56 | 2.48
| UNIPOLAR AMPLITUDE | 1.7 |  | 1.68 | 1.68 | 1.76
| CURRENT | 22 |  | 19.1 | 18.9 | 19.19
|  | RESULT |  | PASS | PASS | PASS
TEST 2: | SENSITIVITY TEST
| BIPOLAR | 1.5 |  | 2 | 2.05 | 1.8
| BIPOLAR | 2.5 |  | 2.7 | 2.8 | 2.5
| BIPOLAR | 3.5 |  | 3.5 | 3.6 | 3.15
| UNIPOLAR | 1.5 |  | 1.95 | 1.8 | 1.8
| UNIPOLAR | 2.5 |  | 2.7 | 2.6 | 2.4
| UNIPOLAR | 3.5 |  | 3.4 | 3.6 | 3.15
| CURRENT | 22 |  | 18.69 | 19.06 | 19.26
|  | RESULT |  | PASS | PASS | PASS
TEST 3: | URGENT TEST
| URGENT AMPLITUDE - UNIPOLAR | 5 |  | 4.88 | 4.88 | 4.72
| URGENT RATE | 60 |  | 60 | 60 | 60
| URGENT PULSE WIDTH | 0.5 |  | 0.497 | 0.493 | 0.498
| URGENT CURRENT | 22 |  | 19.02 | 18.96 | 19.01
|  | RESULT |  | PASS | PASS | PASS
TEST 4: | RATE TEST
| BIPOLAR RATE | 50 |  | 50 | 50 | 50
| LOW ENERGY IMPEDANCE - BIPOLAR | 500 |  | 514 | 527 | 507
| UNIPOLAR RATE | 50 |  | 50 | 50 | 50
| LOW ENERGY IMPEDANCE - UNIPOLAR | 500 |  | 514 | 527 | 514
| CURRENT(LOW ENERGY) | 10 |  | 10.11 | 10.07 | 10.01
|  | RESULT |  | PASS | PASS | PASS
TEST 5: | PULSE WIDTH TEST
| BIPOLAR PULSE WIDTH | 0.3 |  | 0.31 | 0.31 | 0.32
| BIPOLAR PULSE WIDTH | 1.2 |  | 1.19 | 1.19 | 1.2
| BIPOLAR PULSE WIDTH | 1.5 |  | 1.49 | 1.49 | 1.5
| UNIPOLAR PULSE WIDTH | 0.3 |  | 0.32 | 0.31 | 0.32
| UNIPOLAR PULSE WIDTH | 1.2 |  | 1.2 | 1.19 | 1.2
| UNIPOLAR PULSE WIDTH | 1.5 |  | 1.5 | 1.49 | 1.5
| CURRENT | 22 |  | 19.31 | 19.21 | 18.85
|  | RESULT |  | PASS | PASS | PASS
TEST 6: | IMPEDANCE TEST
| BIPOLAR IMPEDANCE | 350 |  | 355 | 359 | 343
| BIPOLAR IMPEDANCE | 1000 |  | 932 | 951 | 932
| BIPOLAR IMPEDANCE | 500 |  | 514 | 520 | 507
| BIPOLAR IMPEDANCE | SHORT |  | <<250 | <<250 | <<250
| BIPOLAR IMPEDANCE | OPEN |  | >>1500 | >>1500 | >>1500
| UNIPOLAR IMPEDANCE | 350 |  | 355 | 359 | 351
| UNIPOLAR IMPEDANCE | 1000 |  | 932 | 951 | 915
| UNIPOLAR IMPEDANCE | 500 |  | 514 | 520 | 514
| UNIPOLAR IMPEDANCE | SHORT |  | <<250 | <<250 | <<250
| UNIPOLAR IMPEDANCE | OPEN |  | >>1500 | >>1500 | >>1500
| CURRENT | 22 |  | 18.97 | 19.12 | 19.78
|  | RESULT |  | PASS | PASS | PASS
TEST 7: | BATTERY TEST
| EOL | BATTERY VOLT EOL |  | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL
| ERI status during EOL | ERI/EOL/BOL |  | EOL | EOL | EOL
| ERI | BATTERY VOLT ERI |  | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI
| EOL status during ERI | ERI/EOL/BOL |  | ERI | ERI | ERI
| BATTERY VOLTAGE (BOL) | BATTERY VOLT BOL |  | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL
| ERI/EOL status during BOL | ERI/EOL/BOL |  | BOL | BOL | BOL
| CURRENT BOL | 22 |  | 20.8 | 21.11 | 21.66
|  | RESULT |  | PASS | PASS | PASS
TEST 8: | MAGNET TEST
| AMPLITUDE BEFORE MAGNET TEST | 3 |  | 3.04 | 2.96 | 2.96
| RATE BEFORE MAGNET TEST | 108 |  | 107.92 | 107.84 | 107.91
| AMPLITUDE DURING MAGNET TEST | 4.5 |  | 4.16 | 4.24 | 4.24
| CURRENT DURING MAGNET TEST | 27 |  | 27.54 | 26.83 | 27.09
| RATE DURING MAGNET TEST (BOL) | 96 |  | 95.82 | 95.87 | 95.82
| RATE ERI MAGNET | 83 |  | 83.34 | 83.37 | 83.43
| RATE EOL MAGNET | 64 |  | 63.96 | 63.98 | 64.01
| AMPLITUDE AFTER MAGNET TEST | 3 |  | 2.96 | 2.96 | 2.96
| RATE AFTER MAGNET TEST | 108 |  | 107.92 | 107.84 | 107.98
| MAGNET TEST (VOO) | PACE COUNT |  | 5 | 5 | 5
| CURRENT | 22 |  | 19.24 | 19.42 | 19.6
|  | RESULT |  | PASS | PASS | PASS
TEST 9: | NOISE TEST
| 11HZ BURST PERIOD TEST - BIPOLAR | Pace Count |  | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - BIPOLAR | Pace Count |  | 0 | 0 | 0
| 11HZ BURST PERIOD TEST - UNIPOLAR | Pace Count |  | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - UNIPOLAR | Pace Count |  | 0 | 0 | 0
| CURRENT | 22 |  | 19.27 | 19.06 | 18.83
|  | RESULT |  | PASS | PASS | PASS
TEST 10: | HIGH ENERGY TEST
| RATE | 120 |  | 119.75 | 119.65 | 119.74
| AMPLITUDE | 7.5 |  | 6.24 | 6.32 | 6.4
| PULSE WIDTH | 1.0 |  | 1.02 | 1.01 | 1.02
| HIGH ENERGY IMPEDANCE | 500 |  | 501 | 514 | 494
| CURRENT | 113 |  | 55.87 | 85.02 | 53.84
|  | RESULT |  | PASS | PASS | PASS
TEST 11: | TRAILING EDGE AND SLEW RATE TEST
| LEADING AMPLITUDE - 5V - VVI - BIPOLAR | 5.3 |  | 5.44 | 5.36 | 5.36
| TRAILING AMPLITUDE - 5V - VVI - BIPOLAR | 4.3 |  | 4.56 | 4.4 | 4.48
| AVERAGE AMPLITUDE - 5V - VVI - BIPOLAR | 4.8 |  | 5 | 4.88 | 4.92
| SLEW RATE - 5V - VVI - BIPOLAR | 0.5-3.0 |  | 1.33 | 1.6 | 2.13
| LEADING AMPLITUDE - 5V - VVI - UNIPOLAR | 5.3 |  | 5.36 | 5.28 | 5.28
| TRAILING AMPLITUDE - 5V - VVI - UNIPOLAR | 4.3 |  | 4.48 | 4.48 | 4.4
| AVERAGE AMPLITUDE - 5V - VVI - UNIPOLAR | 4.8 |  | 4.92 | 4.88 | 4.84
| SLEW RATE - 5V - VVI - UNIPOLAR | 0.5-3.0 |  | 2.13 | 2.13 | 2.13
| CURRENT | 22 |  | 20.3 | 20.41 | 20.48
|  | RESULT |  | PASS | PASS | PASS
TESTED BY:
NAME:
SIGNATURE: