## Sheet
10 | SHREE PACETRONIX LTD.
ADDRESS: PLOT NO. 15, SECTOR II, PITHAMPUR, INDORE –454775 (M.P.) INDIA
TITLE: AUTO TESTER ELECTRICAL TESTING FORM FOR SINGLE CHAMBER WITH R MODE PROMRI CIRCUIT
FORMAT NUMBER – ABCDEF ¦ DEPARTMENT – DEPT. ¦ EFFECTIVE DATE – START DATE ¦ REV – XX
Batch no. | BATCH NO.
Pacemaker Model no. | MODEL NO.
|  |  | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No. | Serial No.
|  | Set Value: | Tolerance: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs: | Obs:
TEST 1: | AMPLITUDE AND LEAKAGE TEST
| BIPOLAR AMPLITUDE IN UNIPOLAR OUTPUT | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| BIPOLAR AMPLITUDE | 3.8 |  | 3.52 | 3.52 | 3.52 | 3.6 | 3.52 | 3.6 | 3.68 | 3.44 | 3.52 | 3.6
| BIPOLAR AMPLITUDE | 2.5 |  | 2.4 | 2.4 | 2.48 | 2.48 | 2.4 | 2.48 | 2.56 | 2.32 | 2.4 | 2.48
| BIPOLAR AMPLITUDE | 1.7 |  | 1.68 | 1.76 | 1.84 | 1.76 | 1.84 | 1.84 | 1.84 | 1.84 | 1.76 | 1.76
| UNIPOLAR AMPLITUDE IN BIPOLAR OUTPUT | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| UNIPOLAR AMPLITUDE | 3.8 |  | 3.44 | 3.6 | 3.6 | 3.52 | 3.52 | 3.52 | 3.6 | 3.44 | 3.52 | 3.68
| UNIPOLAR AMPLITUDE | 2.5 |  | 2.48 | 2.48 | 2.48 | 2.48 | 2.48 | 2.4 | 2.48 | 2.48 | 2.48 | 2.48
| UNIPOLAR AMPLITUDE | 1.7 |  | 1.84 | 1.76 | 1.76 | 1.84 | 1.76 | 1.76 | 1.84 | 1.76 | 1.76 | 1.76
| CURRENT | 22 |  | 18.86 | 17.33 | 18.2 | 17.85 | 17.74 | 17.06 | 16.84 | 17.65 | 17.09 | 17.35
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 2: | SENSITIVITY TEST
| BIPOLAR | 1.5 |  | 1.95 | 2.2 | 2.25 | 2 | 2 | 2.05 | 1.95 | 2.2 | 2 | 2
| BIPOLAR | 2.5 |  | 2.7 | 3 | 2.7 | 3.2 | 2.8 | 2.9 | 2.7 | 3.1 | 3.3 | 2.8
| BIPOLAR | 3.5 |  | 3.4 | 3.8 | 4 | 3.6 | 3.6 | 3.7 | 3.4 | 3.9 | 4.2 | 3.55
| UNIPOLAR | 1.5 |  | 1.9 | 2.2 | 2.2 | 1.85 | 2.05 | 2.1 | 1.8 | 2.2 | Out of range - (High sensitivity) | 2.05
| UNIPOLAR | 2.5 |  | 2.7 | 3.1 | 2.8 | 3 | 2.8 | 2.9 | 2.5 | 2.9 | 3.2 | 2.7
| UNIPOLAR | 3.5 |  | 3.3 | 3.5 | 3.7 | 4 | 3.55 | 3.4 | 1.85 | 3.8 | 4.1 | 3.55
| CURRENT | 22 |  | 17.53 | 17.36 | 18.07 | 17.77 | 17.95 | 17.64 | 17.06 | 18.3 | 18.17 | 17.46
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | FAIL | PASS
TEST 3: | URGENT TEST
| URGENT AMPLITUDE - UNIPOLAR | 5 |  | 4.72 | 4.96 | 4.72 | 4.72 | 4.8 | 4.64 | 4.8 | 4.72 | 4.88 | 4.88
| URGENT RATE | 60 |  | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60
| URGENT PULSE WIDTH | 0.5 |  | 0.483 | 0.491 | 0.484 | 0.484 | 0.487 | 0.483 | 0.487 | 0.485 | 0.487 | 0.483
| URGENT CURRENT | 22 |  | 17.29 | 17.69 | 18.74 | 18.13 | 17.9 | 17.19 | 17.63 | 17.92 | 17.33 | 17.33
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 4: | RATE TEST
| BIPOLAR RATE | 50 |  | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
| LOW ENERGY IMPEDANCE - BIPOLAR | 500 |  | 447 | 426 | 432 | 437 | 442 | 447 | 442 | 432 | 447 | 442
| UNIPOLAR RATE | 50 |  | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
| LOW ENERGY IMPEDANCE - UNIPOLAR | 500 |  | 447 | 426 | 432 | 437 | 437 | 453 | 437 | 432 | 442 | 442
| CURRENT(LOW ENERGY) | 10 |  | 9.83 | 8.7 | 9.86 | 9.39 | 10.03 | 9.59 | 8.69 | 9.35 | 9.67 | 9.14
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 5: | PULSE WIDTH TEST
| BIPOLAR PULSE WIDTH | 0.3 |  | 0.3 | 0.31 | 0.3 | 0.3 | 0.31 | 0.3 | 0.31 | 0.3 | 0.31 | 0.3
| BIPOLAR PULSE WIDTH | 1.2 |  | 1.18 | 1.19 | 1.19 | 1.19 | 1.19 | 1.18 | 1.19 | 1.19 | 1.19 | 1.18
| BIPOLAR PULSE WIDTH | 1.5 |  | 1.49 | 1.5 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49
| UNIPOLAR PULSE WIDTH | 0.3 |  | 0.3 | 0.31 | 0.3 | 0.3 | 0.31 | 0.3 | 0.31 | 0.3 | 0.31 | 0.3
| UNIPOLAR PULSE WIDTH | 1.2 |  | 1.18 | 1.19 | 1.19 | 1.19 | 1.19 | 1.18 | 1.19 | 1.19 | 1.19 | 1.18
| UNIPOLAR PULSE WIDTH | 1.5 |  | 1.49 | 1.5 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49 | 1.49
| CURRENT | 22 |  | 18.02 | 17.86 | 18.83 | 18.09 | 18.38 | 20.68 | 17.54 | 17.6 | 18.33 | 18.07
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 6: | IMPEDANCE TEST
| BIPOLAR IMPEDANCE | 350 |  | 375 | 363 | 363 | 371 | 371 | 379 | 371 | 371 | 379 | 375
| BIPOLAR IMPEDANCE | 1000 |  | 881 | 865 | 865 | 865 | 865 | 881 | 881 | 881 | 881 | 1157
| BIPOLAR IMPEDANCE | 500 |  | 437 | 426 | 426 | 432 | 432 | 442 | 437 | 437 | 437 | 437
| BIPOLAR IMPEDANCE | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| BIPOLAR IMPEDANCE | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| UNIPOLAR IMPEDANCE | 350 |  | 375 | 375 | 363 | 375 | 367 | 384 | 371 | 371 | 375 | 375
| UNIPOLAR IMPEDANCE | 1000 |  | 881 | 881 | 865 | 865 | 865 | 881 | 881 | 881 | 881 | 881
| UNIPOLAR IMPEDANCE | 500 |  | 437 | 442 | 426 | 437 | 426 | 442 | 437 | 437 | 437 | 432
| UNIPOLAR IMPEDANCE | SHORT |  | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250 | <<250
| UNIPOLAR IMPEDANCE | OPEN |  | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500 | >>1500
| CURRENT | 22 |  | 18.17 | 18.11 | 18.03 | 18.67 | 18.02 | 18.03 | 17.22 | 18.91 | 17.36 | 18.31
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 7: | BATTERY TEST
| EOL | BATTERY VOLT EOL |  | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL | BATTERY VOLT EOL
| ERI status during EOL | ERI/EOL/BOL |  | EOL | EOL | EOL | EOL | EOL | EOL | EOL | EOL | EOL | EOL
| ERI | BATTERY VOLT ERI |  | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI | BATTERY VOLT ERI
| EOL status during ERI | ERI/EOL/BOL |  | ERI | ERI | ERI | ERI | ERI | ERI | ERI | ERI | ERI | ERI
| BATTERY VOLTAGE (BOL) | BATTERY VOLT BOL |  | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL | BATTERY VOLT BOL
| ERI/EOL status during BOL | ERI/EOL/BOL |  | BOL | BOL | BOL | BOL | BOL | BOL | BOL | BOL | BOL | BOL
| CURRENT BOL | 22 |  | 19.71 | 19.52 | 21.99 | 19.66 | 20.73 | 18.83 | 18.74 | 21.03 | 20.42 | 20.71
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 8: | MAGNET TEST
| AMPLITUDE BEFORE MAGNET TEST | 3 |  | 2.96 | 3.04 | 2.96 | 2.8 | 2.96 | 2.96 | 2.96 | 2.96 | 2.88 | 2.88
| RATE BEFORE MAGNET TEST | 108 |  | 107.99 | 107.88 | 107.99 | 107.98 | 107.92 | 107.98 | 107.93 | 107.98 | 107.98 | 108.07
| AMPLITUDE DURING MAGNET TEST | 4.5 |  | 4.16 | 4.24 | 4.08 | 4.08 | 4.08 | 4.08 | 4.24 | 4.16 | 4.08 | 4.16
| CURRENT DURING MAGNET TEST | 27 |  | 21.86 | 25.35 | 26.53 | 25.33 | 25.69 | 21.65 | 21.34 | 21.95 | 25.21 | 25
| RATE DURING MAGNET TEST (BOL) | 96 |  | 95.9 | 95.91 | 95.86 | 95.82 | 95.82 | 95.82 | 95.91 | 95.86 | 95.82 | 95.94
| RATE ERI MAGNET | 83 |  | 83.4 | 83.41 | 83.37 | 83.34 | 83.34 | 83.43 | 83.41 | 83.45 | 83.43 | 83.43
| RATE EOL MAGNET | 64 |  | 63.99 | 63.99 | 63.92 | 63.96 | 63.96 | 64.01 | 63.94 | 63.97 | 64.01 | 64.01
| AMPLITUDE AFTER MAGNET TEST | 3 |  | 2.88 | 2.96 | 2.96 | 2.88 | 3.04 | 2.8 | 2.88 | 2.96 | 2.88 | 2.88
| RATE AFTER MAGNET TEST | 108 |  | 107.88 | 107.93 | 107.99 | 107.92 | 107.92 | 107.98 | 107.89 | 107.88 | 107.98 | 108.07
| MAGNET TEST (VOO) | PACE COUNT |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| CURRENT | 22 |  | 18.98 | 18.22 | 19.77 | 18.74 | 19.21 | 18.69 | 18.45 | 18.77 | 18.8 | 17.93
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 9: | NOISE TEST
| 11HZ BURST PERIOD TEST - BIPOLAR | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - BIPOLAR | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 11HZ BURST PERIOD TEST - UNIPOLAR | Pace Count |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 5HZ BURST PERIOD TEST - UNIPOLAR | Pace Count |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| CURRENT | 22 |  | 17.51 | 17.12 | 18.55 | 18.25 | 18.52 | 17.53 | 17.04 | 18.12 | 17.37 | 17.47
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 10: | HIGH ENERGY TEST
| RATE | 120 |  | 119.75 | 119.7 | 119.7 | 119.69 | 119.75 | 119.82 | 119.7 | 119.74 | 119.83 | 119.92
| AMPLITUDE | 7.5 |  | 6.24 | 6.4 | 6.32 | 6.4 | 6.32 | 6.4 | 6.4 | 6.24 | 6.32 | 6.4
| PULSE WIDTH | 1.0 |  | 1.01 | 1.01 | 1.01 | 1.01 | 1.01 | 1.01 | 1.01 | 1.01 | 1.01 | 1.01
| HIGH ENERGY IMPEDANCE | 500 |  | 437 | 412 | 426 | 432 | 421 | 442 | 432 | 421 | 437 | 426
| CURRENT | 113 |  | 79.93 | 88.2 | 102.53 | 98.86 | 78.41 | 78.67 | 82.2 | 81.31 | 102.88 | 88.57
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 11: | TRAILING EDGE AND SLEW RATE TEST
| LEADING AMPLITUDE - 5V - VVI - BIPOLAR | 5.3 |  | 5.2 | 5.36 | 5.28 | 5.28 | 5.28 | 5.28 | 5.2 | 5.2 | 5.28 | 5.36
| TRAILING AMPLITUDE - 5V - VVI - BIPOLAR | 4.3 |  | 4.08 | 4.48 | 4.32 | 4.16 | 4.32 | 4.24 | 4.24 | 4.32 | 4.32 | 4.48
| AVERAGE AMPLITUDE - 5V - VVI - BIPOLAR | 4.8 |  | 4.64 | 4.92 | 4.8 | 4.72 | 4.8 | 4.76 | 4.72 | 4.76 | 4.8 | 4.92
| SLEW RATE - 5V - VVI - BIPOLAR | 0.5-3.0 |  | 1.6 | 1.87 | 2.13 | 1.87 | 2.4 | 1.87 | 1.87 | 1.33 | 2.4 | 1.6
| LEADING AMPLITUDE - 5V - VVI - UNIPOLAR | 5.3 |  | 5.12 | 5.28 | 5.28 | 5.12 | 5.28 | 5.2 | 5.28 | 5.2 | 5.2 | 5.44
| TRAILING AMPLITUDE - 5V - VVI - UNIPOLAR | 4.3 |  | 4.24 | 4.48 | 4.4 | 4.24 | 4.32 | 4.24 | 4.4 | 4.48 | 4.32 | 4.32
| AVERAGE AMPLITUDE - 5V - VVI - UNIPOLAR | 4.8 |  | 4.68 | 4.88 | 4.84 | 4.68 | 4.8 | 4.72 | 4.84 | 4.84 | 4.76 | 4.88
| SLEW RATE - 5V - VVI - UNIPOLAR | 0.5-3.0 |  | 2.13 | 2.13 | 2.4 | 2.13 | 2.13 | 2.13 | 2.13 | 1.6 | 2.13 | 1.87
| CURRENT | 22 |  | 19.98 | 19.4 | 20.1 | 19.59 | 20.29 | 19.22 | 19.07 | 19.89 | 19.54 | 18.87
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TEST 12: | RATE RESPONSIVE RATE TEST
| VVIR MODE CURRENT | 27 |  | 23.03 | 22.4 | 24.46 | 23.12 | 24.88 | 22.27 | 23.22 | 23.78 | 23.65 | 22.97
| RATE BEFORE ACTIVITY | 60 |  | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60
| RATE CHANGE AFTER ACTIVITY |  |  | 86.01 | 81.47 | 72.61 | 77.31 | 66.05 | 80.75 | 79.72 | 94.53 | 69.7 | 79.82
| R MODE CURRENT AFTER ACTIVITY | 35 |  | 31.33 | 29.38 | 27.58 | 28.13 | 25.45 | 28.97 | 27.98 | 32.67 | 26.27 | 27.45
|  | RESULT |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
TESTED BY:
NAME:
SIGNATURE: